## Slide 1
lh
D
A starting position for calibration is when there is no mirror (horizontal beam), the center of the detector is at the sample height and the detector face is vertical. At this point Hu=Hd=0, X = D. We need to determine X, Hu, and Hd as functions of D and 2θ.
2θ
lv
Hd
Hu
L
X
x
y
z
dH
lh
Hd=0
Hu=0
X = D
X = 0

## Slide 2
lh
D
X = (D + lh) cos(2θ) – lv sin(2θ) – lh
Hu = (D + lh) sin(2θ) + lv cos(2θ) – lv
Hd = Hu + L tan(2θ) + dH/cos(2θ) – dH
2θ
lv
Hu
X
D
2θ
2θ
Hd
Hu
L
L
x
y
z
dH
dH
lh
lv
lh
L

## Slide 3
BM MEDM slide
D
2θ
lv
X
L
sample
x
y
z
CCD
lh
dH
lh
Hu= 0
Hd
Hu
Hd= 0

## Slide 4
Inverse motion equations
Direct equations:
Inverse equations (we presume that Hd must move in sync with Hu and
denote Du = D + lh ):
Du 2cos2(2θ) – 2 Du lv sin(2θ) cos(2θ) + lv2 sin2(2θ) = (X + lh)2
Du 2sin2(2θ) + 2 Du lv sin(2θ) cos(2θ) + lv2 cos2(2θ) = (Hu + lv)2
Du 2 = (X + lh)2 + (Hu + lv)2 – lv2  D =SQRT[(X + lh)2 + (Hu + lv)2 – lv2] – lh
+
X = Du cos(2θ) – lv sin(2θ) – lh
Hu = Du sin(2θ) + lv cos(2θ) – lv
Du cos(2θ) – lv sin(2θ) = X + lh
Du sin(2θ) + lv cos(2θ) = Hu+ lv
(1)
Then, from (1):
X = (D + lh) cos(2θ) – lv sin(2θ) – lh
Hu = (D + lh) sin(2θ) + lv cos(2θ) – lv
Hd = Hu + L tan(2θ) + dH/cos(2θ) – dH
sin(2θ) =
Hu (D + lh) + (D – X) lv
(D + lh) 2 + lv2

## Slide 5
dX = d(D) cos(2θ) – d(2θ) [Du sin(2θ) + lv cos(2θ)]
dHu = d(D) sin(2θ) + d(2θ) [Du cos(2θ) – lv sin(2θ)]
dHd = d(Hu) + d(2θ) [L+dH sin(2θ)] / cos2(2θ)
Differential equations
Direct equations ( we denote: Du = D + lh ):
Inverse equations (we presume that Hd must move in sync with Hu):
d(D) = [dX(X + lh ) + dHu(Hu + lv)] / (D + lh )
d(2θ) = [d(D)cos(2θ) – dX] / (Hu + lv) --- from Eqs. for dX and Hu